|  |  | # Environmental Health and Safety # STANDARD OPERATING GUIDELINE # AUTOCLAVE VALIDATION |  |  |
| --- | --- | --- | --- | --- |
| # Section: | # Environmental Health and Safety Department # Environmental Control |  | # Date of Issue: |  |
|  |  |  | # Issued By: | # Environmental Health &Safety |
| # Part: | # Environmental Management System - Environmental Guideline |  | # Revision #: |  |
|  |  |  | # Revision Date: |  |
| # Pages: | # 3 |  | # Revised By: |  |

Purpose/Scope

  - This procedure is used to assess the efficiency and performance of autoclaves.

Responsibilities

  - Each operating college/department has the primary responsibility for assuring proper use and validation of autoclave equipment in order to ensure efficiency of sterilization procedures.

References

  - Biosign Steam-24 Biological Indicators Product Specification

Definitions

  - Autoclave Validation: A quality assurance procedure used to ensure that the autoclave reaches adequate temperature for an adequate amount of holding time to sterilize biological agents and wastes.

  - Biological Indicators: Will demonstrate that adequate temperature and holding time has been achieved in an autoclave; should be used during validation to ensure efficient sterilization.  The most commonly used biological indicator is Bacillus stearothermophius, being the most resistant to steam autoclaving.

Procedure

  - Each autoclave should be validated every 40 operating hours.  See the “Cumulative Time” on the Autoclave Operation Log to determine when validation is necessary.

  - Each time the autoclave is validated, the date of validation, pass/fail and other relevant information should be noted on the Autoclave Validation Log.

  - When the cumulative time reaches 38-40 hours, a validation test must be performed.  Wait for the next run of the autoclave, and follow the procedures as follows.

    - Be sure to wear gloves, a lab coat, and any other necessary personal protective equipment (PPE) when working with biological agents and waste.

    - Reference supplier instruction for the Getinge Biosign Steam-24 Biological Indicator.

    - Make sure the Getinge Biosign Steam-24 indicators being used are not expired, and record the lot # and expiration date on the Autoclave Validation Log.

    - Place Getinge Biosign Steam-24 indicator in the center of the load being sterilized/under the bag being autoclaved.

    - Use another sample of the Getinge Biosign Steam-24 indicator not autoclaved as a control.

    - Process load as per standard operating procedure.

      - When autoclave cycle is complete, remove and incubate the indicator spores and the positive control as per supplier’s instructions.  The Biosign Steam-24 Indicator should be incubated at 55 C ±3 C for 24 hours.

      - After the 24 hour incubation period, look for a color change in the indicator media.

        - A color change from Red to Yellow indicates positive growth (failure).

        - No color change indicates no growth (passing).

    - Check the positive control (sample not autoclaved) for color change to ensure the validity of results.

      - If the test cannot be considered valid, the validation should be performed again with new indicators.

      - If the test is determined valid, record the results and test parameters on the Autoclave Operation Log and Autoclave Validation Log.

      - If the test fails, record the results and test parameters on the Autoclave Operation Log and Autoclave Validation Log and notify the area PI immediately.

      - Label autoclave “Out of Service” until a passing validation test is achieved.

    - Dispose of the used indicators and contaminated PPE as biological waste.

    - Attachments

      - Autoclave Operation Log

      - Autoclave Validation Log

- Ehs on 'finsrvx6'\health and safety programs\osha programs\biosafety\rit institute biosafety committee\autoclave validation procedure april 2013.docx